## Calibratin_raw_data pH
2-Point-Calibration pH accordig to common pH-Meters (Küster-Thiel (1982)) | [17 empty] | t [°C] | 0,025 M KH2PO4 + 0,025M Na2HPO4 | 0,008695M KH2PO4+0,03043 Na2HPO4
| Micrometrics-Vietnam | [6 empty] | Date: | 2004-03-24 | Temperature correktion | [7 empty] | 0 | 6.984 | 7.534
2004-07-05 | Sensor no.:2 |  |  | S1 | S2 |  |  | slope S1 intercept S1 | slope S2 intercept S2 | pH 4.00 ±0.02 |  | pH 7.00 ±0.02 |  | pH 9.21 ±0.02 |  | Interpolation pH(T) Mettler-Toledo Buffers |  | 5 | 6.951 | 7.5
Buffer 2 and 1 | T[°C] Glass-Elektrode | T[°C] Ag/Ag+ Elektrode | pH(t) | signal [V] | signal [V] | Standard/Buffer |  | -56.6616 | -56.6616 | 0 | 4.05 | 0 | 7.13 | 0 | 9.52 | T [°C] | 22.6 | 10 | 6.923 | 7.472
pH(t) 4.00 | 24 | 24 | 4.01 | 0.166 | 0.166 | Mettler-Toledo |  | 394.2324 | 394.2324 | 5 | 4.04 | 5 | 7.07 | 5 | 9.45 | pH9.21 | 9.2332 | 15 | 6.9 | 7.448
pH(t) 7.00 | 24 | 24 | 7 | 0 | 0 | Mettler-Toledo |  | slope recomended 100% - 85% |  | 10 | 4.02 | 10 | 7.05 | 10 | 9.38 | T [°C] | 50 | 20 | 6.881 | 7.429
pH(t) 9.00 | 24 | 24 | 9.21 | -0.129 | -0.129 | Mettler-Toledo |  | -58.9614 | -50.1172 | 15 | 4.01 | 15 | 7.02 | 15 | 9.32 | pH7.00 | 6.938 | 25 | 6.865 | 7.413
Tavg[°C]: | 24 |  | EXAMPLES |  |  |  | Null correktion | a = Null S1 (pH7.00) | a = Null S2 (pH7.00) | 20 | 4 | 20 | 7 | 20 | 9.26 | T [°C] | 22.6 | 30 | 6.853 | 7.4
ID | T[°C] Glass-Elektrode | T[°C] Ag/Ag+ Elektrode | measured signal1[V] | measured signal2[V] | pH(S1) | pH(S2) | S1[V] | 0 | 0 | 25 | 3.99 | 25 | 6.99 | 25 | 9.21 | pH4.00 | 3.9948 | 35 | 6.844 | 7.389
sample_01pH_01 | 25 |  | 0.166 | 0 | 6.9551 | 6.9298 | 0 | Fscorr. = S1 | Fscorr. = S2 | 30 | 3.99 | 30 | 6.98 | 30 | 9.16 | Reference temperature |  | 38 | 6.84 | 7.384
sample_02pH_02 | 22.6 |  | 0.1535 | 0.1453 | 4.2317 | 4.3771 | S2[V] | 1.0406 | 1.0406 | 35 | 3.99 | 35 | 6.97 | 35 | 9.11 | y=-0,66057x+224,243 |  | 40 | 6.838 | 7.38
sample_03pH_03 | 22.6 |  | -0.125 | 0.1 | 9.1701 | 5.1803 | 0 |  |  | 40 | 3.99 | 40 | 6.96 | 40 | 9.06 |  |  | 45 | 6.834 | 7.373
|  |  |  | slope check (>85% of theoretical value) |  |  |  | 96,1% OK | 96,1% OK | 45 | 3.99 | 45 | 6.95 | 45 | 9.03 |  |  | 50 | 6.833 | 7.367
|  |  |  | null check (±0.030 V of theoretical value at pH7.00) |  |  |  | 0,000 OK | 0,000 OK | 50 | 4 | 50 | 6.94 | 50 | 8.99
|  | 59.1598 |  |  |  | 7.01 | 0.01
|  | 61.561 |  |  |  | 7
Messbereich ADS1115_4096 |  | 6.9577 |  |  |  | 6.99 | [14 empty] | 5039.7483
20241122 | pH7 | 2.5102 |  | pH4,01 | 3.0085 | 0.4983 | 8.3607
«< | [9 empty] | TABELLE 1 | [6 empty] | Formeln zur Temperaturkompensation der in situ Messungen und der Kalibration (Tabelle1):
[10 empty] | Bei einwertigen Ionen wie dem Wasserstoffion gelten | [6 empty] | pH t2 = pH t1 + 0.0114(t1-t2) ; t1 = Kalibriertemperatur , t2 = in situ Temperatur
[10 empty] | folgende Temperaturabhängigkeiten der Nernstspannung. | [10 empty] | 7.7 Calculation of the in situ pH for seawater (Grasshoff)
[10 empty] | J [°C] | S [mV] | J [°C] | S [mV] | J [°C] | S [mV] | [5 empty] | 7.7.1 Temperature correction
[10 empty] | 0 | 54.2 | 35 | 61.14 | 70 | 68.08 | [5 empty] | The method described is based on a pH measurement at constant temperature. The in situ
[10 empty] | 5 | 55.19 | 40 | 62.13 | 75 | 69.08 | [5 empty] | value of the pH must then be computed. Gieskes (1969a, 1970) has shown that for practical
[10 empty] | 10 | 56.18 | 45 | 63.12 | 80 | 70.07 | [5 empty] | purposes one temperature coefficient can be usod because of the fact that the determinations
[10 empty] | 15 | 57.17 | 50 | 64.12 | 85 | 71.06 | [5 empty] | of the equilibrium constants in the carbonic acid system, the boric acid and different ion pair
[10 empty] | 20 | 58.16 | 55 | 65.11 | 90 | 72.05 | [5 empty] | formations involved do not justify an apparent accuracy of the pH data of better than ± 0.02
[10 empty] | 25 | 59.16 | 60 | 66.1 | 95 | 73.04 | [5 empty] | pH units.
[10 empty] | 30 | 60.15 | 65 | 67.09 | 100 | 74.04 | [5 empty] | If the data of Buch (1951) and those of Lyman (1956) are corrected for the NBS pH scale, a
[10 empty] | A change of 1 pH unit causes a change of voltage: | [10 empty] | temperature coefficient of 0.0114 seems to be representative (Gieskes, 1969a, 1970). The pH
[10 empty] | 55,19 mV (5°C), 59,16 mV (25°C) or 63,12 mV (45°C). All Nernst linear regression lines crossing the | [10 empty] | at the in situ temperature can then be calculated from
[10 empty] | Isotherme at pH = 7.00. | [10 empty] | PHt2 = pHt1 + 0.0114 (t1—t2)
[10 empty] | That means: The pH is temperature dependend. A calibration is needed. | [10 empty] | where
t [°C] | 0,025 M KH2PO4 + 0,025M Na2HPO4 | 0,008695M KH2PO4+0,03043 Na2HPO4 | [18 empty] | t1 = measuring temperature
0 | 6.984 | 7.534 |  | 6.8992e-05 | -0.0063854 | 6.9817 | #N/A | #N/A |  | Potentials measured against the Standard Hydrogen Elektrode at 25°C, were reactands are in the concentration of 1 mol/L | [10 empty] | t2 = in situ temperature
5 | 6.951 | 7.5 |  | 2.0449e-06 | 0.00010533 | 0.0011409 | #N/A | #N/A |  | are Normal Voltages or Standard potentials. Redox pairs, | [10 empty] | In view of the uncertainty of the temperature coettilcient of ±0.0010, the accuracy of the
10 | 6.923 | 7.472 |  | 0.99927 | 0.0015054 | #N/A | #N/A | #N/A |  | which deliver electrons to the NHE, have a negative sign, Redox pairs, which | [10 empty] | correction will be ± 0.02 pH units if t1 and t2 differ by 20 °C.
15 | 6.9 | 7.448 |  | 6179.9653 | 9 | #N/A | #N/A | #N/A |  | are taking electrons getting a positive sign.
20 | 6.881 | 7.429 |  | 0.028009 | 2.0395e-05 | #N/A | #N/A | #N/A |  | When you sort the redox pairs in order to there negative voltage against NHE
25 | 6.865 | 7.413 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | we get the voltage series oder redox series. At the beginning of the voltage series, | [11 empty] | Standard potentials for Reference elektrodes in dependence of temperature against the Normal hydrogen elektrode NHE [mV] acc. to Dr. W. Ingold KG - Frankfurt
30 | 6.853 | 7.4 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | are Redox pairs with the highest tendency to spend elektrones ; | [12 empty] | Calomel | Ag/AgCl |  | Regression coeffizients for Reference elektrode types:
35 | 6.844 | 7.389 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | At the end of the voltage series, are redox pairs, which are taking electrons. | [11 empty] | T [°C] | Elektrolyte: sat KCl-sol. | Elektrolyte: sat. KCl-sol. | Elektrolyte: 3M KCl-sol. | 1 |  | 2 |  | 3
38 | 6.84 | 7.384 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | pH- Meter (a Voltmeter): the pH- Meter is using the phenomenon of potential differences dE at membranes with different pHs | [11 empty] | Elektrode | 1 | 2 | 3 | a | b | a | b | a | b
40 | 6.838 | 7.38 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | [H+] concentrations (pH1, pH2) in contact : | [11 empty] | 0 | 260.2 | 227.6 | 223.6 | -0.63857 | 260.3929 | -0.78629 | 228.0214 | -0.66057 | 224.2429
45 | 6.834 | 7.373 |  | #N/A | #N/A | #N/A | #N/A | #N/A |  | delta E [V] = - 2,303 RT / F (pH2 - pH1) | [11 empty] | 10 | 254.2 | 220.4 | 218.1 | 0.0067763 | 0.13122 | 0.016106 | 0.31189 | 0.041781 | 0.80908
50 | 6.833 | 7.367 | [7 empty] | calculation of dE pro pH-Eiheit zu berechnen ergibt sich: | [11 empty] | 15 | 250.9 | 216.6 | 215 | 0.99955 | 0.16366 | 0.99832 | 0.389 | 0.98425 | 1.0091
[10 empty] | delta E [V] = - 2,303 RT / F (pH2 - pH1) / (pH2-pH1) | [11 empty] | 20 | 247.7 | 212.6 | 211.8 | 8880.4 | 4 | 2383.2913 | 4 | 249.9693 | 4
[10 empty] | Basic equation is the NERNST equation: | [11 empty] | 25 | 244.4 | 208.2 | 206.2 | 237.8679 | 0.10714 | 360.643 | 0.60529 | 254.5402 | 4.0731
[10 empty] | E = E0 - (RT / nF) ln [red] / [ox] | [11 empty] | 30 | 241.1 | 204.1 | 204.7
[22 empty] | mV for T [°C]: | 247.6214 | 212.2957 | 205.7469
[8 empty] | 45 |  | If pH1 is known, pH2 can be calculated from delta E. Each pH- Meter have to be calibrated before use. (> Festlegung von | [11 empty] | 28
[8 empty] | 6.8341 |  | pH1. The pH meter is an precise instrument, which needs maintenance.
[10 empty] | The "Elektrode" hold wet every time. 3m KCl storing solution. | [18 empty] | y=-0,66057x+224,243
The inner Buffer is also 3M KCl. Die Glas membrane is
[10 empty] | The electrode is very sensitive against against mechanical and chemical influences. pH- Meter are inprecise over pH 12 (> Alkali error) | [11 empty] | Küster Thiel, Rechentafeln für die chemische Analytik 102. Auflage, Walter de Gruyter, Berlin-New York, 1982
pH 7 is according to a signal of 0 mV.
The difference to 0 mV and the pH7-Signal is called Asymetry.
Standard potentials for Reference elektrodes in dependence of temperature against the Normal hydrogen elektrode NHE [mV]
| Mettler-Toledo 7,00 |  |  | Mettler-Toledo 4,01 | [6 empty] | Calomel (Eo =0,2690V) | Ag/AgCl (Eo= 0,2222V) |  | Regressions coeffizients for Reference elektrode types:
| 5 |  |  | 5 | [5 empty] | T [°C] | Elektrolyte: sat. KCl-sol. | Elektrolyte: sat. KCl-sol. | Elektrolyte: 3M KCl-sol. | 1 |  | 2 |  | 3
| 10 |  |  | 10 | [5 empty] | Elektrode | 1 | 2 | 3 | a | b | a | b | a | b
| 15 |  |  | 15 | [5 empty] | 0 | 260.2 | 227.6 | 223.6 | -0.63857 | 260.3929 | -0.78629 | 228.0214 | -0.66057 | 224.2429
| 20 |  |  | 20 | [5 empty] | 10 | 254.2 | 220.4 | 218.1 | 0.0067763 | 0.13122 | 0.016106 | 0.31189 | 0.041781 | 0.80908
| 25 | 7 |  | 25 | 4.01 |  |  |  |  | 15 | 250.9 | 216.6 | 215 | 0.99955 | 0.16366 | 0.99832 | 0.389 | 0.98425 | 1.0091
| 30 |  |  | 30 | [5 empty] | 20 | 247.7 | 212.6 | 211.8 | 8880.4 | 4 | 2383.2913 | 4 | 249.9693 | 4
| 35 |  |  | 35 | [5 empty] | 25 | 244.4 | 208.2 | 206.2 | 237.8679 | 0.10714 | 360.643 | 0.60529 | 254.5402 | 4.0731
| 40 |  |  | 40 | [5 empty] | 30 | 241.1 | 204.1 | 204.7
| 45 |  |  | 45 | [5 empty] | U [mV] for T [°C]: | 244.4286 | 208.3643 | 207.7286
25

## Messung Kalibration 20250117
[8 empty] | Theoretischer Wert (mV) pro pH-Einheit | Pro 3 pH-Einheiten
[8 empty] | 58.1677 | 174.5031
Sonde | Puffer (pH) | Bezugsspannung [V] | Messwert [Volt] | T [°C] | calc. Ratio | mV/dpH |  |  | pH 7 | pH 4 | Messwert(Probe) mV
Leonie | 4 | 3 | 2.0491 | 20 | 2049.125 | 181.8295 | Slope | -0.017442 | 1533.125 | 2049.125 | 2049
Leonie | 7.02 | 3 | 1.5331 | 20 | 1533.125 |  | Intercept pH | 7.1926
|  |  | 1.774 | 20 |  |  | pH(Probe) | 4.0007
| Puffer (pH) | Bezugsspannung [V] | Messwert [Volt] | T [°C] | calc. Ratio |  |  |  | pH 7 | pH 4 | Messwert(Probe) mV
Matthias | 4 | 3 | 3.0556 | 20 | 3055.6 |  | Slope | -0.017428 | 2539.2 | 3055.6 | 2539
Matthias | 7.02 | 3 | 2.5392 | 20 | 2539.2 |  | Intercept pH | 13.0372
|  |  | 1.774 | 20 |  |  | pH(Probe) | 7.0012
[7 empty] | Quelle: | https://github.com/DFRobot/DFRobot_PH/blob/master/python/raspberrypi/DFRobot_PH.py