## Dissolved Oxygen
Temperature [°C]	Voltage [V]	Lineare Regression		t [°C]	O2 [%]	U [V]			T[K]	Salinität	100% O2-Sättigung Wert [mg/L] nach Benson & Krause					Messung 2		U[V]	O2[mg/L]
16,125	1,61025	0,0466607142857143	0,863685267857143	21,5	100	1,865	0,01829	0,0359999999999999	294,65	0	8,82861492785682					BEG Zentralkläranlage		2,0085	18,4
15,125	1,570875	0,00151543925066265	0,020809430617138	21,5	0	0,036	0	0	294,65	0						2025-02-18		1,80125	16,6
14,125	1,5305	0,995798486620948	0,00633953721835096				1	0								VE Wasser mit Natriumsulfit		1,747	16,2
13,125	1,47375	948,037906137188	4				#NV	0										1,654	14,9
12,125	1,434375	0,0381013895089286	0,000160758928571428				1,6726205	0								ROH		1,5395	14
11,125	1,376875			Leitungswasser	U [V]	O2 [%]										DO		1,451625	13,1
				21,5	1,3	35,1111111111112										18.4		0,37225	1,7
																Analog value: 16067		0,108875	0,3
																Voltage: 2.0085V

																16.2
																Analog value: 13975
																Voltage: 1.747V

																14.9
																Analog value: 13232
																Voltage: 1.654V

																14
																Analog value: 12316
																Voltage: 1.5395V

																13.1
																Analog value: 11613
																Voltage: 1.451625V

																0.3
																Analog value: 872
																Voltage: 0.108875V

																1.7
																Analog value: 2978
																Voltage: 0.37225V

																16.6
																Analog value: 14409
	Benson und Krause Formel															Voltage: 1.80125V
Signal(100%O2) vs Temperatur
Temperaturkompensation
t [°C]
U [V]
Kalibration Oxygen
0 - 100% bei 21.5°C
c(O2) [%]
U [V]
Gelöster Sauerstoff vs Spannung
U[V]
O2[mg/L]

## TDS
Messung 1		U[V]	EC [µS/m]					Messung 2		U[V]	EC[µS/cm]
BEG Zentralkläranlage		0,007875	0					BEG Zentralkläranlage		0,066625	48
2025-02-04		1,294875	1346					2025-02-18		0,13325	101
0,01 mol/L Kaliumchlorid		1,7375	1276					0,01 mol/L Kaliumchlorid		0,256125	186
und VE Wasser					0,339625	277
				0,470125	375
				0,54775	440
ROH:					ROH:		1,594875	1400

destilliert:					Leitffähigkeit in Mikrosiemens
Analog value: 56					0.01mil/L kcl verdünnt
Voltage: 0.007875V					48 Mikrosiemens
				Analog value: 525
mit kalibrationslösung					Voltage: 0.066625V
1276 Mikrosiemens
Analog value: 13904					101 Mikrosiemens
Voltage: 1.7375V					og value: 1058
temp: 20,5					Voltage: 0.13325V

1346 Mikrosiemens					186 Mikrosiemens
Analog value: 10362					Analog value: 2043
Voltage: 1.294875V					Voltage: 0.256125V

				277 Mikrosiemens
				Analog value: 2712
				Voltage: 0.339625V

				375 Mikrosiemens
				Analog value: 3756
				Voltage: 0.470125V

				440 Mikrosiemens
				Analog value: 4380
				Voltage: 0.54775V

				1400 Mikrosiemens
				Analog value: 12767
				Voltage: 1.594875V
Leitfähigkeit in Abhängigkeit von Spannung Messungen 1
U[V]
EC[µS/m]
Leitfähigkeit in Abhängigkeit von Spannung
U[V]
EC[µS/m]

## Turbidity
	Messung 1		U[V]	Turbidity[FNU]	Art				Messung 2		U[mV]	Turb[FNU]	U[V]	Turb[FNU]
	BEG Zentralkläranlage		2,069375	0,24	VE Wasser				BEG Zentralkläranlage		1930	96	1,93	96
	2025-02-04		2,043875	41	Kläranlagen Probe				2025-02-18		1874	193	1,874	193
			2,025375	41	Kläranlagen Probe				Gemessen mit ADC.py		1683	400	1,683	400
			1,685625	555	Bariumchlorid				12 bit ADC		1653	500	1,653	500
			0,804	???	Mikroorganismen						1573	530	1,573	530
									„Reaktionswasser“		1442	850	1,442	850
	mit einem FNU messgerät								Abfall Produkt Emery
	0.24 fnu								850 fnu
	Analog value: 16553								1442
	Voltage: 2.069375V								530 fnu
									1573
	41 fnu								500 fnu
	Analog value: 16352								1653
	Voltage: 2.043875V								400 fnu
	gain 2/3								1683
	Analog value: 10803								193 fnu
	Voltage: 2.025375V								1874
									96 fnu
	bariumchlorid								1930
	555fnu
	Analog value: 8994
	Voltage: 1.685625V

	mikroorganismen
gain	Analog value: 3732
1	Voltage: 0.804V
1
2/3
2/3
2/3
FNU
volt
fnu
Trübheit
Turbidity[FNU]
U[V]
Turbidity vs Spannung
U[V]
Turbidity[FNU]

## pH
				Theoretischer Wert (mV) pro pH-Einheit	Pro 3 pH-Einheiten
				58,17	174,50
Sonde	Puffer (pH)	Bezugsspannung [V]	Messwert [Volt]	T [°C]	calc. Ratio	mV/dpH			pH 7	pH 4	Messwert(Probe) mV
Leonie	4,00	3	2,049125	20	2049,125	181,829470	Slope	-0,0175581395348837	1533,125	2049,125	2049
Leonie	7,02	3	1,533125	20	1533,125		Intercept pH	7,21
			1,774	20			pH(Probe)	4,00073158914729

	Puffer (pH)	Bezugsspannung [V]	Messwert [Volt]	T [°C]	calc. Ratio				pH 7	pH 4	Messwert(Probe) mV
Matthias	4,00	3	3,0556	20	3055,6		Slope	-0,017428350116189	2539,2	3055,6	2539
Matthias	7,02	3	2,5392	20	2539,2		Intercept pH	13,04
			1,774	20			pH(Probe)	7,00116189000775

				Quelle:	https://github.com/DFRobot/DFRobot_PH/blob/master/python/raspberrypi/DFRobot_PH.py

## pH beispiel Matthias
2-Point-Calibration pH accordig to common pH-Meters (Küster-Thiel (1982))					t [°C]	0,025 M KH2PO4 + 0,025M Na2HPO4	0,008695M KH2PO4+0,03043 Na2HPO4
	Micrometrics-Vietnam					Date:	24.03.2004	Temperature correktion					0	6,984	7,534
05.07.2004	Sensor no.:2			S1	S2			slope S1 intercept S1	slope S2 intercept S2	pH 4.00 ±0.02		pH 7.00 ±0.02		pH 9.21 ±0.02		Interpolation pH(T) Mettler-Toledo Buffers		5	6,951	7,500
Buffer 2 and 1	T[°C] Glass-Elektrode	T[°C] Ag/Ag+ Elektrode	pH(t)	signal [V]	signal [V]	Standard/Buffer		-56,661577	-56,661577	0	4,05	0	7,13	0	9,52	T [°C]	22,6	10	6,923	7,472
pH(t) 4.00	24	24	4,01	0,1660	0,1660	Mettler-Toledo		394,2323599	394,2323599	5	4,04	5	7,07	5	9,45	pH9.21	9,23	15	6,9	7,448
pH(t) 7.00	24	24	7,00	0,0000	0,0000	Mettler-Toledo		slope recomended 100% - 85%		10	4,02	10	7,05	10	9,38	T [°C]	50	20	6,881	7,429
pH(t) 9.00	24	24	9,21	-0,1290	-0,1290	Mettler-Toledo		-58,961396	-50,117187	15	4,01	15	7,02	15	9,32	pH7.00	6,94	25	6,865	7,413
Tavg[°C]:	24,0		EXAMPLES				Null correktion	a = Null S1 (pH7.00)	a = Null S2 (pH7.00)	20	4,00	20	7,00	20	9,26	T [°C]	22,6	30	6,853	7,400
ID	T[°C] Glass-Elektrode	T[°C] Ag/Ag+ Elektrode	measured signal1[V]	measured signal2[V]	pH(S1)	pH(S2)	S1[V]	0,0000	0,0000	25	3,99	25	6,99	25	9,21	pH4.00	3,99	35	6,844	7,389
sample_01pH_01	25,0		0,1660	0,0000	6,955	6,93	0,0000	Fscorr. = S1	Fscorr. = S2	30	3,99	30	6,98	30	9,16	Reference temperature		38	6,84	7,384
sample_02pH_02	22,6		0,1535	0,1453	4,23	4,38	S2[V]	1,0406	1,0406	35	3,99	35	6,97	35	9,11	y=-0,66057x+224,243		40	6,838	7,380
sample_03pH_03	22,6		-0,1250	0,1000	9,17	5,18	0,0000			40	3,99	40	6,96	40	9,06			45	6,834	7,373
				slope check (>85% of theoretical value)				96,1% OK	96,1% OK	45	3,99	45	6,95	45	9,03			50	6,833	7,367
				null check (±0.030 V of theoretical value at pH7.00)				0,000 OK	0,000 OK	50	4,00	50	6,94	50	8,99

		59,1598193062306				7,01	0,01
		61,56				7,00
Messbereich ADS1115_4096		6,9577				6,99					5039,7483096087
20241122	pH7	2,5102		pH4,01	3,0085	0,4983	8,36073825503356
pH vs T [°C] of Mettler-Toledo commercial available technical Buffers / T [°C] / pH
«<					TABELLE 1					Formeln zur Temperaturkompensation der in situ Messungen und der Kalibration (Tabelle1):
				Bei einwertigen Ionen wie dem Wasserstoffion gelten					pH t2 = pH t1 + 0.0114(t1-t2) ; t1 = Kalibriertemperatur , t2 = in situ Temperatur
				folgende Temperaturabhängigkeiten der Nernstspannung.					7.7 Calculation of the in situ pH for seawater (Grasshoff)
				J [°C]	S [mV]	J [°C]	S [mV]	J [°C]	S [mV]					7.7.1 Temperature correction
				0,0	54,2	35,0	61,14	70,0	68,08					The method described is based on a pH measurement at constant temperature. The in situ
				5,0	55,19	40,0	62,13	75,0	69,08					value of the pH must then be computed. Gieskes (1969a, 1970) has shown that for practical
				10,0	56,18	45,0	63,12	80,0	70,07					purposes one temperature coefficient can be usod because of the fact that the determinations
				15,0	57,17	50,0	64,12	85,0	71,06					of the equilibrium constants in the carbonic acid system, the boric acid and different ion pair
				20,0	58,16	55,0	65,11	90,0	72,05					formations involved do not justify an apparent accuracy of the pH data of better than ± 0.02
				25,0	59,16	60,0	66,1	95,0	73,04					pH units.
				30,0	60,15	65,0	67,09	100,0	74,04					If the data of Buch (1951) and those of Lyman (1956) are corrected for the NBS pH scale, a
				A change of 1 pH unit causes a change of voltage:					temperature coefficient of 0.0114 seems to be representative (Gieskes, 1969a, 1970). The pH
				55,19 mV (5°C), 59,16 mV (25°C) or 63,12 mV (45°C). All Nernst linear regression lines crossing the					at the in situ temperature can then be calculated from
				Isotherme at pH = 7.00.					PHt2 = pHt1 + 0.0114 (t1—t2)
				That means: The pH is temperature dependend. A calibration is needed.					where
t [°C]	0,025 M KH2PO4 + 0,025M Na2HPO4	0,008695M KH2PO4+0,03043 Na2HPO4					t1 = measuring temperature
0	6,984	7,534		6,89916846151111E-05	-0,00638537898439459	6,98174622209257	#NV	#NV		Potentials measured against the Standard Hydrogen Elektrode at 25°C, were reactands are in the concentration of 1 mol/L					t2 = in situ temperature
5	6,951	7,500		2,04487082219307E-06	0,00010533453027911	0,00114087976037728	#NV	#NV		are Normal Voltages or Standard potentials. Redox pairs,					In view of the uncertainty of the temperature coettilcient of ±0.0010, the accuracy of the
10	6,923	7,472		0,999272370402052	0,00150536872881657	#NV	#NV	#NV		which deliver electrons to the NHE, have a negative sign, Redox pairs, which					correction will be ± 0.02 pH units if t1 and t2 differ by 20 °C.
15	6,9	7,448		6179,96530032468	9	#NV	#NV	#NV		are taking electrons getting a positive sign.
20	6,881	7,429		0,0280092714515794	2,03952150872894E-05	#NV	#NV	#NV		When you sort the redox pairs in order to there negative voltage against NHE
25	6,865	7,413		#NV	#NV	#NV	#NV	#NV		we get the voltage series oder redox series. At the beginning of the voltage series,					Standard potentials for Reference elektrodes in dependence of temperature against the Normal hydrogen elektrode NHE [mV] acc. to Dr. W. Ingold KG - Frankfurt
30	6,853	7,400		#NV	#NV	#NV	#NV	#NV		are Redox pairs with the highest tendency to spend elektrones ;					Calomel	Ag/AgCl		Regression coeffizients for Reference elektrode types:
35	6,844	7,389		#NV	#NV	#NV	#NV	#NV		At the end of the voltage series, are redox pairs, which are taking electrons.					T [°C]	Elektrolyte: sat KCl-sol.	Elektrolyte: sat. KCl-sol.	Elektrolyte: 3M KCl-sol.	1		2		3
38	6,84	7,384		#NV	#NV	#NV	#NV	#NV		pH- Meter (a Voltmeter): the pH- Meter is using the phenomenon of potential differences dE at membranes with different pHs					Elektrode	1	2	3	a	b	a	b	a	b
40	6,838	7,380		#NV	#NV	#NV	#NV	#NV		[H+] concentrations (pH1, pH2) in contact :					0	260,2	227,6	223,6	-0,638571428571433	260,392857142857	-0,786285714285714	228,021428571429	-0,660571428571432	224,242857142857
45	6,834	7,373		#NV	#NV	#NV	#NV	#NV		delta E [V] = - 2,303 RT / F (pH2 - pH1)					10	254,2	220,4	218,1	0,00677630927178959	0,13122266479196	0,0161061529629303	0,311894310984492	0,0417807689381282	0,809081111445553
50	6,833	7,367					calculation of dE pro pH-Eiheit zu berechnen ergibt sich:					15	250,9	216,6	215,0	0,999549772635181	0,163663417676999	0,99832446088069	0,389000550862631	0,9842500662876	1,00910143904651
				delta E [V] = - 2,303 RT / F (pH2 - pH1) / (pH2-pH1)					20	247,7	212,6	211,8	8880,39999999945	4	2383,29132247657	4	249,969322858207	4
				Basic equation is the NERNST equation:					25	244,4	208,2	206,2	237,867857142857	0,107142857142864	360,643047619048	0,605285714285722	254,540190476191	4,07314285714293
				E = E0 - (RT / nF) ln [red] / [ox]					30	241,1	204,1	204,7
				mV for T [°C]:	247,6	212,3	205,7
				45		If pH1 is known, pH2 can be calculated from delta E. Each pH- Meter have to be calibrated before use. (> Festlegung von					28,0
				6,834126		pH1. The pH meter is an precise instrument, which needs maintenance.
				The "Elektrode" hold wet every time. 3m KCl storing solution.					y=-0,66057x+224,243
				The inner Buffer is also 3M KCl. Die Glas membrane is

				The electrode is very sensitive against against mechanical and chemical influences. pH- Meter are inprecise over pH 12 (> Alkali error)					Küster Thiel, Rechentafeln für die chemische Analytik 102. Auflage, Walter de Gruyter, Berlin-New York, 1982
				pH 7 is according to a signal of 0 mV.
				The difference to 0 mV and the pH7-Signal is called Asymetry.

				Standard potentials for Reference elektrodes in dependence of temperature against the Normal hydrogen elektrode NHE [mV]
	Mettler-Toledo 7,00			Mettler-Toledo 4,01					Calomel (Eo =0,2690V)	Ag/AgCl (Eo= 0,2222V)		Regressions coeffizients for Reference elektrode types:
	5			5					T [°C]	Elektrolyte: sat. KCl-sol.	Elektrolyte: sat. KCl-sol.	Elektrolyte: 3M KCl-sol.	1		2		3
	10			10					Elektrode	1	2	3	a	b	a	b	a	b
	15			15					0	260,2	227,6	223,6	-0,638571428571428	260,392857142857	-0,786285714285715	228,021428571429	-0,660571428571429	224,242857142857
	20			20					10	254,2	220,4	218,1	0,00677630927178941	0,131222664791957	0,0161061529629303	0,311894310984492	0,0417807689381281	0,809081111445551
	25	7,00		25	4,01					15	250,9	216,6	215,0	0,999549772635181	0,163663417676995	0,99832446088069	0,389000550862631	0,9842500662876	1,0091014390465
	30			30					20	247,7	212,6	211,8	8880,39999999992	4	2383,29132247656	4	249,969322858209	4
	35			35					25	244,4	208,2	206,2	237,867857142857	0,107142857142858	360,643047619048	0,605285714285723	254,540190476191	4,07314285714291
	40			40					30	241,1	204,1	204,7
	45			45					U [mV] for T [°C]:	244,4	208,4	207,7
				25,0